## Slide 1
High Throughput Sequencing
(NEXT GENERATION SEQUENCING) Technologies and Applications
Petrus Tang, Ph.D. (鄧致剛)
Graduate Institute of Basic Medical Sciences
and
Bioinformatics Center, Chang Gung University.
petang@mail.cgu.edu.tw
http://petang.cgu.edu.tw

## Slide 2
James Watson and Francis Crick with their DNA model at the Cavendish Laboratories in 1953

## Slide 3
Do You Read Me?

## Slide 4
Sanger Dideoxy Sequencing
Sanger, F. et al. Nature 24, 687–695 (1977).
Sanger, F., Nicklen, S. & Coulson, A.R. Proc. Natl. Acad. Sci. USA 74, 5463–5467 (1977).

## Slide 5
Chain-termination sequencing
Chemical sequencing
Recombinant DNA

## Slide 6
Nature, 15 February 2001
Vol. 409, Pages 813-960
Science, 16 February 2001
Vol. 291, Pages 1145-1434
April 2003: High-Resolution Human Genome
February 2001: Completion of the Draft Human Genome
Nature, 23 April 2003
Vol. 422, Pages 1-13
>10 years to finish
USD 3 billion

## Slide 7
3,000,000,000 bp

## Slide 8
ABI 3730 XL DNA Sequencer
96/384 DNA sequencing in 2 hrs, approximately 600-1000 readable bps per run.
1-4 MB bps/day
A human genome of 3GB need 750 days to finish

## Slide 9
Whitehead Institute

## Slide 10
273 authors
254 authors

## Slide 13
Next Generation Sequencing Technology
Massively Parallel Signature Sequencing (MPSS)

## Slide 14
NEXT GENERATION SEQUENCING

## Slide 15
Next Generation Sequencing Technology
Massively Parallel Signature Sequencing (MPSS)
(Genome Sequencer, GS)
(Genome Analyzer, GA)
(Sequencing by Oligonucleotide Ligation
and Detection, SOLiD)

## Slide 16
NATURE METHODS | VOL.5 NO.1 | JANUARY 2008
Nature Methods' Method of the Year 2007 goes to next-generation sequencing. This series of articles showcase how these novel sequencing methods came into their own in 2007 and the incredible impact they promise to have in a variety of research applications. The Methods to Watch feature provide a glimpse and a wish list for future Methods of the Year.
Next Generation Sequencing Technology
SOLID
http://marketing.appliedbiosystems.com/
SOLEXA
http://www.illumina.com/pages.ilmn?ID=250
454 GS FLX
http://www.454.com/
Massively Parallel Signature Sequencing (MPSS)

## Slide 17
Throughput of NGS machines (2007-2009)
Throughput = total number of reads x read length
High throughput
Short read/run
Illumina
SOLiD
Low throughput-
Long read/run
Roche
(~4Gb/4 days)

## Slide 18
Throughput of NGS machines: Roche (2012)

## Slide 19
Throughput of NGS machines: Roche

## Slide 20
Throughput of NGS machines: SOLiD 4 (2013)
In 2014 ABI SOLiD was replaced by the Ion Torrent sequencing platforms
(Proton and Personal Genome Machine)

## Slide 21
NATURE METHODS | VOL.5 NO.1 | JANUARY 2008
Nature Methods' Method of the Year 2007 goes to next-generation sequencing. This series of articles showcase how these novel sequencing methods came into their own in 2007 and the incredible impact they promise to have in a variety of research applications. The Methods to Watch feature provide a glimpse and a wish list for future Methods of the Year.
Next Generation Sequencing Technology
SOLID
http://marketing.appliedbiosystems.com/
SOLEXA
http://www.illumina.com/pages.ilmn?ID=250
454 GS FLX
http://www.454.com/
Massively Parallel Signature Sequencing (MPSS)
X
X

## Slide 22
2014 NGS Machines: Hiseq X

## Slide 23
2015 NGS Machines: Hiseq 3000/4000

## Slide 24
2016 NGS Machine: NovaSeq6000

## Slide 25
Throughput of NGS machines ~2016
High throughput
short read
Illumina
Low throughput,
Very long read
PacBio
X
X
X
X
NovaSeq6000

## Slide 26
NEXT GENERATION SEQUENCING
Basis Principles
Temple fragmentation (library) Amplification Sequencing Detection

## Slide 27
Illumina Sequencing by Synthesis (SBS)

## Slide 28
Solexa Genome Analyzer

## Slide 29
Clusters Array
Template Bridging
Cluster Generation

## Slide 30
Denaturation
Ready For Sequencing

## Slide 31
Sequencing-By-Synthesis

## Slide 32
Incorporation
Fluorescent Detection
Sequence Generated At Every Site On The Array

## Slide 33
From Debbie Nickerson, Department of Genome Sciences, University of Washington, http://tinyurl.com/6zbzh4

## Slide 34
From Debbie Nickerson, Department of Genome Sciences, University of Washington, http://tinyurl.com/6zbzh4

## Slide 35
Illumina sequencing-by-synthesis
Sample preparation.
Fragments of DNA are ligated to end adapters, denatured and bound at one end to a solid surface already coated with a dense layer of the adapters. Each single stranded fragment is immobilized at one end, while its free end ‘bends over’ and hybridizes to a complementary adapter on the surface, which initiates the synthesis of the complementary strand in the presence of amplification reagents. Multiple cycles of this solid-phase amplification followed by denaturation create clusters of ~1,000 copies of single-stranded DNA molecules distributed randomly on the surface.
Sequencing with reversible terminators.
Synthesis reagents, added to the flow cell, consist of primers, DNA polymerase and four differently labeled, reversible terminator nucleotides. After incorporation of a nucleotide, which is identified by its color, the 3′ terminator on the base and the fluorophore are removed, and the cycle is repeated for a read length of 30–35 nucleotides.

## Slide 36
Combinatorial Probe-Anchor Ligation (cPAL) sequencing
Complete Genomics was founded in 2006 and in March 2013 was acquired by
BGI-Shenzhen for USD 117.6 million.

## Slide 37
BGISEQ-50
BGISEQ-500

## Slide 38
Sanger dideoxy-sequencing
Massive parallel sequencing
Bead-based em-PCR and sequencing by ligation
Massive parallel sequencing and
single molecule sequencing
Single molecule sequencing and nanopore technology?
ABI 3730XL
Roche 454 FLX, Life Technologies SOLiD,
Illumina Hiseq, Life Tech Ion Torrent, BGIseq
Dover Systems' Polonator
Oxford Nanopore Technologies,, Electron Optica, Electronic BioSciences, GenapSys, Genia, GnuBIO, IBM, LaserGen, Lightspeed Genomics, NABsys, NobleGen Biosciences, QuantuMDx, Reveo, Stratos Genomics, Two Pore Guys, ZS Genetics………………….
Pacific Biosciences (single-molecule real-time DNA sequencing (SMRT) technology)
Helicos (true single-molecule-sequencing (tSMS) technology)
VisiGen Biotechnologies (real-time, single-molecule sequencing
fluorescence resonance energy transfer (FRET) technology)
Evolution of Sequencing Technology

## Slide 39
NEW GENERATION SEQUENCING
Basis Principles
Temple fragmentation (library) Amplification Sequencing Detection
NEXT-NEXT-NEXT-

## Slide 40
Single Molecule Real-Time Sequencing
VisGen Biotechnologies
Pacific Biosciences
Helicos Biosciences
Real-time monitoring of PCR activity
Read-out by fluorescence resonance energy transfer between polymerase and nucleotides or
Waveguides allow direct observation of polymerase and fluorescently labeled nucleotides
Helicos BioSciences said in a regulatory filing that it has sold certain patent applications covering molecular diagnostic applications of its technology to Sequenom for $1.3 million. The patent applications cover methods for detecting fetal nucleic acids and diagnosing fetal abnormalities. Helicos also warned that if it is "unable to successfully raise additional capital, we may have to cease operations and/or seek bankruptcy protection."

## Slide 41
Single Molecule Real-Time (SMRT) Sequencing

## Slide 42
Sequel & RSII systems
Sequel
RSII

## Slide 44
Oxford Nanopore Sequencing
Nucleic acids driven through a nanopore.
Differences in conductance of pore provide readout.
This diagram shows a protein nanopore set in an electrically resistant membrane bilayer. An ionic current is passed through the nanopore by setting a voltage across this membrane. If an analyte passes through the pore or near its aperture, this event creates a characteristic disruption in current. By measuring that current it is possible to identify the molecule in question. For example, this system can be used to distinguish the four standard DNA bases and G, A, T and C, and also modified bases. It can be used to identify target proteins, small molecules, or to gain rich molecular information for example to distinguish the enantiomers of ibuprofen or molecular binding dynamics.

## Slide 47
2014 E. coli genome
Avg. read length: 5.4 kb
Longest read: 10 kb
2015
Mean read length: 2 kb
Median read length : 1 kb
Longest read: 98 kb
R9.4 flow cell
Throughput: 5-10 Gb

## Slide 52
SmidgION: nanopore sensing for use with mobile devices

## Slide 54
Cost of NGS machines
High throughput, short read/run
COST USD800,000
Low throughput-
long read/run
COST USD1,000,000

## Slide 55
Cost of NGS machines
High throughput, short read/run
COST USD800,000
Medium throughput, short read/run
COST USD200,000-400,000
Low throughput-
long read/run
COST USD1,000,000

## Slide 56
Personal Sequencer
Ion Torrent/Ion Proton
Personal Genome Machine
Miseq/MiniSeq
Personal Sequencing System

## Slide 60
314 chip
316 chip
Ion Torrent Sequencing

## Slide 62
(2012)

## Slide 63
Miniseq (2016)

## Slide 66
GeneReader
Early-access customers in 2013,
Commercial launch in early 2014.
20 flowcells configuration
20-400M reads per flow cell
12 minutes per base sequencing-by-synthesis (50-bases ~10 hours)

## Slide 67
Droplet-based Sequencing

## Slide 68
Microfluidics, the science that deals with the behavior, precise control and manipulation of fluids that are geometrically constrained to a small, typically sub-millimeter scale, emerged in the beginning of the 1980s and is used widely in technologies such as inkjet print heads, DNA chips, lab-on-a-chip technology, micro-propulsion and micro-thermal technologies. GnuBIO has adapted technology developed in Dr. David Weitz laboratory at Harvard. This technology is combined with proprietary molecular biology assays to create a unique next-generation DNA sequencing technology.
The science behind this GnuBIO sequencing technology utilizes droplet microfluidics to perform the biochemical reactions for sequencing inside of tiny picoliter -sized aqueous drops. Each droplet acts as a discrete reaction vessel, like a miniature test tube, where the sequencing assay is performed.
Droplet-based Sequencing

## Slide 69
Storage & Computing Power
Next gen sequencers generated Giga bp to Tera bp of data

## Slide 70
http://allseq.com/knowledgebank

## Slide 71
NEXT GENERATION SEQUENCING Applications & Analysis

## Slide 72
Applications on Biomedical Sciences

## Slide 73
High-throughput Sequencing (HTseq) Platforms

## Slide 74
Applications on Biomedical Sciences